Spindle—a New Type of differential Pressure Flowmeter

Xiao Ming

Nanjing Youyang Control Technology Ltd, Qingshuiting West Road, Nanjing 211102, China

SPINDLE is a new type of differential-pressure flowmeter that generates the pressure drop via a spindle-like central body as the primary flow element. The fluid flows through the annular channel between the central body and the wall of the pipe. As the central body is forcing an annular channel flow, SPINDLE has a powerful ability to condition the different flow characteristics upstream of the meter, as present in industrial environments, into a very stable, repeatable and well characterized flow, which provides a constant performance with high accuracy. Different from other primary flow elements, such as orifice plate, Venturi shaped cone etc, SPINDLE minimizes the drag of the central body as its shape has been optimized by numerical calculations, so that no separation occurs, while keeping the differential pressure as high as possible. The streamlined shape also provides a self-cleaning function. The SPINDLE flow meter has been calibrated in different facilities, including the water calibration rig of NMI of China and the air calibration rig of Jiangsu Province Calibration Center. The calibration at PTB in Germany with different piping configurations for flow disturbances according to OIML R 137 has verified the high performance of SPINDLE at configurations as typical for industry. The calibrations have shown that the SPINDLE flow meter exhibits a very constant dependence of discharge coefficient C on the Reynolds number Re for Re larger than 1 x 105. Uncertainties of 0.2 % in water and 0.5 % in air have been achieved because of the very stable flow and the very rigid structure of SPINDLE. The repeatability of the measurement of the flow rate is better than 0.05 % for water.