The influence of ultrasonic transducer temperature adaptability on the measurement results

Xiaoyu Li ,  Jun Mu , Dunli Liu

Xinjiang Uygur Autonomous Region Research Institute of Measurement and Testing

No.188 Hebei Esat Road,Urumqi,China

E-mail：562185424@qq.com

# Abstract

Portable ultrasonic flowmeter used to measure the fluid flowmeters, it take the advantage of the propagation characteristics of the ultrasonic.With its advantage of  touch free measurement, wide measuring range, ease to install, simper operation of testing, it’s is considered to be a better measuring device in reducing runoff.It’s widely used in electrical power, petroleum, chemical industry, especially the water supply system.In the process of using ultrasonic flowmeter to measure the flow, the accuracy of the measurement results is impacted to some extent due to the difference of the transducer temperature adaptability. By analyzing the experimental data set, this article discusses the impact of the measuring result caused by the improper use of the transducer..

Key word: Portable ultrasonic flowmeter ;  ultrasonic transducer  temperature ; the measuring results

# 1. Introduction

Portable ultrasonic flowmeter used to measure the fluid flowmeters,it take the advantage of the propagation characteristics of the ultrasonic.With its advantage of touch free measurement,wide measuring range, ease to install, simper operation of testing,it’s is considered to be a better measuring device in reducing runoff.It’s widely used in electrical power, petroleum,hemical industry, especially the water supply system. has been developing rapidly in recent years. This article discusses the change of temperature effect on the results of the measurement conclusion have realistic meaning for research to improve measurement accuracy..

# 2. The working principle of the Portable ultrasonic flowmeter

Portable ultrasonic flowmeter belongs to the method of time difference ultrasonic flowmeter: Portable ultrasonic flowmeter using pairs of acoustic transmitter and acoustic receiver (transducer), When an acoustic wave propagation in the fluid, along the flow direction of sound waves is accelerated. but against the direction of fluid motion sound velocity is lowe. They formed between the time difference is proportional to the fluid velocity.

As shown in Figure 1, transducer 1 send ultrasonic to 2, sound velocity c increase speed because the fluid flow velocity vm:

（1）

Figure1 Working principle of portable ultrasonic flowmeter

Whereas, transducer 2 send ultrasonic to 1, sound velocity is slow down because the fluid flow velocity vm:

（2）

L------ The length of the ultrasonic transmission path between the transducer, m;

θ------ Ultrasonic wave incident Angle;

t12------Travel time from the ultrasonic transducer 1 to 2,s;

t21------Travel time from the ultrasonic transducer 2 to 1,s;

vm-----The average velocity of fluid,m/s;

c------The ultrasonic velocity in static fluid,m/s;

Formula 1 minus the formula 2:

=                （3）

# 3. relationship between the ultrasonic wave propagation velocity and temperature

Figures and graphics are often helpful in describing the experimental set up and summarising results.  All graphs, drawings, and photographs must be referenced in text, in the correct numerical sequence.  An example of the figure and caption is shown in Figure 1. The figures should be no wider than the column of text. The ultrasonic velocity in fluid is affected by the density and the temperature of the fluid.For example, in the water when the temperature (0 to 35)℃，Temperature rise per 1 ℃, the sound velocity 4.6 m/s increased, the speed of sound varies with the change in temperature is linear；but when the temperature (35 to 70)℃, sound velocity changing with temperature increase slowly slow performance is nonlinear；When further raise the temperature to 70 ℃ or more,With the temperature rise velocity began to decline, this trend is nonlinear.

Figure 2: Water temperature corresponding to the sound velocity

# 4. The ultrasonic propagation path

When using the portable ultrasonic flowmeter for flow measurement, Transducer emit sound waves through the pipe wall first, then enter the fluid through the pipe wall again,sound waves is received by another transducer,Because sound waves through the different material,the ultrasonic propagation path changed direction three times because of refraction. As shown in Figure 3. The obvious,refracting before and after a path to produce a certain deviation.This path deviation, directly affect the measurement result.

Figure 3 Schematic diagram of ultrasonic propagation path

# 5. The influence of temperature change on the ultrasonic transducer

By reference [1],With the increase of temperature the optimum working frequency of piezoelectric ultrasonic transducer is reduced.When the temperature over 65℃.The optimum working frequency with the increase of temperature drop is more obvious.This phenomenon leads to receive signal is too weak,can not meet the test requirements,this requires transducer can be adjusted according to the working environment temperature working frequency,realize the temperature of the working frequency compensation..

# 6. Conclusion

When using the portable ultrasonic flowmeter:The first,the different temperature range to use different ultrasonic transducer, Especially under the condition of high temperature,It is very important to avoid using normal temperature ultrasonic transducer. Exceed transducer rated temperature range,In severe cases will damage the ultrasonic transducer,Even if there is no damage to the transducer, also can because of the high temperature lead to deformation caused by the ultrasonic propagation path of change,Make Precision measurement Lose their meaning, this is Used to determine the installation position. The second, the velocity of sound change were caused by temperature change  is obvious, the nonlinear change appears especially at high temperature,the sound velocity change trend in different fluid are also different,this requires us to do further research, to corrected the sound velocity,enables the sound velocity of medium temperature corresponding to more accurately.The third,The influence of temperature on the measurement accuracy is embodied in many aspects，So it is very important. In the process of measuring temperature of accurate or not.Usually we will choose the temperature measurement point attachment as medium temperature, the temperature measured by the instrument,installed on the pipeline of the instrument.Usually we will choose the temperature measurement point attachment as medium temperature, the temperature measured by the instrument,it is installed on the pipeline,In some condition on site near point under the condition of temperature can not be provided,We may use temperature value provided by the user,When using these parameters, we should ensure the reliable numerical.

In the process of using Portable ultrasonic flowmeter for flow measurement. The temperature change of changes in the speed of sound, not just the change of the transducer working frequency, fluid pressure, density and other parameters will also change.How to make a convenient flow measurement instrument to play its role, is we need to study hard in the future.

# References

[1] Zheng Xi-bin,Bao Min, Study on Automatic Temperature Compensation System of Operating Frequency of Ultrasonic Piezoelectric Transducer[J],Journal of Zhejiang Sci Tech University, 2015.11(33)835-841